# Screenless Health Band — Request for Quotation

Screen-less health monitoring band · OEM/ODM · private label · Kazakhstan

| Document | RFQ summary — the full requirement list is Annex 1 (98 requirements, sent with this letter) |
| --- | --- |
| Version | 1.0 |
| Date | 2 September 2026 |
| Issued by | A healthcare IT company, Republic of Kazakhstan |
| Pilot order | 1,000 units (also quote 2k / 5k / 10k for the scale-up) |
| Reply deadline | 14 calendar days from receipt |

## 1. What we are buying

We are a healthcare IT company from Kazakhstan. We run a medical information system used by state and private clinics and a patient mobile application with a nationwide user base. We are adding a hardware layer: a branded, screen-less wrist band for remote monitoring of patients with chronic conditions between doctor visits — heart failure, hypertension in patients 65+, COPD/asthma, type 2 diabetes — plus a retail consumer channel.

Three facts that define the whole project:

1. No display, no stock app. All value lives in our own iOS/Android application and in the doctor's panel inside our MIS. We will not ship your app (H Band, Da Fit, VeryFit, GloryFit or similar) to end users.

2. The data is ours. Device → our app → our servers in Kazakhstan. No data may transit or be stored in your cloud, at any point, including activation and firmware updates. This is a legal requirement in our market, not a preference.

3. Two paths run in parallel on the same hardware. Path A — the band is registered in Kazakhstan as a Class 2a medical device with our company as the certificate holder. Path B — the same device sold as a wellness product while registration is in progress. We need a supplier who can support the Path A documentation package without changing the hardware.

Two SKUs. SKU-A — no ECG: PPG (HR/HRV), SpO2, skin temperature, 6-axis IMU, haptics, LED, tap input. This is what we intend to order. SKU-B — SKU-A plus single-lead ECG: quote separately, we want the price and the roadmap now without raising the risk class today.

## 2. The twelve conditions that decide the deal

Each line references the detailed requirement in Annex 1. A "No" on any of these normally ends the evaluation, whatever the price.

| # | Condition | Annex 1 |
| --- | --- | --- |
| 1 | Native BLE SDK for both iOS and Android, documentation available before we order samples (NDA acceptable) | DB-501, DB-503 |
| 2 | Beat-to-beat RR intervals available to our app — not only an averaged BPM | DB-203 |
| 3 | SDK exposes the full data set: HR, SpO2, temperature, activity, sleep stages, respiratory rate, wear status, device events | DB-504, Annex 1 App. A |
| 4 | Fully offline operation — no connection to your cloud at any point, confirmed in writing | DB-506 |
| 5 | Complete BLE protocol specification as an annex to the supply contract — our insurance against SDK discontinuation | DB-507 |
| 6 | No licence, per-device or annual fee for SDK, protocol or API access | DB-508 |
| 7 | Haptic reminder + one-tap patient acknowledgement as a timestamped event delivered to our app — this is the core mechanism of the product, and without a display there is no other way | DB-402, DB-403 |
| 8 | Battery ≥ 7 days on our duty cycle and an on-device buffer of ≥ 5 days of full-resolution data with no phone | DB-302, DB-307 |
| 9 | ISO 13485 held by the manufacturing site, certificate scan supplied. "We can assist the customer" is not ISO 13485 | DB-701 |
| 10 | CE marking plus existing test reports: IEC 60601-1 / -1-2 / -1-11, ISO 80601-2-61 (SpO2), ISO 10993 (skin contact) — state honestly what exists and what must be commissioned | DB-703–DB-706 |
| 11 | You accept the arrangement "we are the legal manufacturer, your site is the manufacturing location" and supply the corresponding technical file in English | DB-710, DB-711 |
| 12 | Full private label — our brand only, on device, strap, charger, packaging, manual, BLE advertising name and firmware strings | DB-801 |

One more, easy to miss. If your platform offers cuffless blood pressure or glucose "trend" features, they must be switchable off in our production firmware and absent from the label, packaging and BLE feature matrix (DB-411). We will not claim what we cannot register.

Accuracy we will verify on samples against a certified reference: heart rate MAE ≤ 5 bpm (or ≤ 10 %), SpO2 Arms ≤ 3.5 % over 70–100 % per ISO 80601-2-61 methodology, skin temperature ± 0.3 °C, ≥ 90 % of beats detected in a clean 5-minute night window.

## 3. Pricing

Quote FOB (state the port) and DDP Almaty or Astana, Kazakhstan, in USD, with the validity period of the quotation.

| Line item | 1,000 pcs (pilot) | 2,000 pcs | 5,000 pcs | 10,000 pcs |
| --- | --- | --- | --- | --- |
| a) Device (SKU-A), unit price |  |  |  |  |
| b) Charger / dock and cable |  |  |  |  |
| c) Custom retail packaging + printed guide (RU/KK) |  |  |  |  |
| d) One-time NRE: mould / ID change / firmware customisation |  |  |  |  |
| e) SDK, app or cloud fees, if any |  |  |  |  |
| f) Certification and test costs to be borne by us |  |  |  |  |
| g) SKU-B (ECG) delta per unit |  |  |  |  |
| Total landed estimate, DDP Kazakhstan |  |  |  |  |

Budget guidance, so neither side wastes time. Our business case is built on a landed cost that keeps a *medical* device cheaper than a consumer fitness band in our market. At volume, a base SKU-A quotation materially above USD 11.5 FOB at 10,000 units is unlikely to proceed, unless the premium is justified by certification coverage that removes cost from our registration budget. We understand that the pilot batch of 1,000 units carries a premium over that figure — state that premium explicitly rather than hiding it, and tell us honestly where your price lands at each tier.

Also state, in one line each: whether you accept a pilot order of 1,000 units and on what terms; your MOQ for OEM and for ODM; payment terms; IP ownership of a customised design, mould and firmware; and whether you would grant exclusivity for our customised model in Kazakhstan and Central Asia (Annex 1, DB-1101–DB-1106).

## 4. What to send us

We are not looking for a catalogue. Suppliers who return the first three items below move straight to the sample stage; brochures without the response matrix will not be evaluated.

1. Completed response matrix — Annex 1, Section 15. All [M] requirements answered explicitly.

2. Completed pricing table — Section 3 above.

3. SDK documentation or repository link (send us your NDA and we will sign it).

4. ISO 13485 and ISO 9001 certificate scans.

5. CE Declaration of Conformity and the test reports of DB-704 / DB-705.

6. ISO 10993 biocompatibility reports with the tested material identified.

7. Datasheet of the proposed model with sensor part numbers, plus photographs or drawings with dimensions.

8. Cost and lead time for 3–5 engineering samples with SDK access.

If your closest existing model already meets most of this specification, say so and give the model number. We prefer a proven platform over a new development, provided the data-access and certification requirements are met.

## 5. How we will decide, and when

| Group | Weight |
| --- | --- |
| Data access — SDK, RR/raw data, offline operation, protocol annex, no fees | 30 % |
| Regulatory readiness — ISO 13485 at the site, existing test reports, willingness to be the manufacturing site under our registration | 25 % |
| Hardware fit — battery, buffer, haptics and acknowledgement input, sensor quality, IP rating | 20 % |
| Commercial — pilot price at 1k and landed cost at 10k, NRE, MOQ, payment terms, exclusivity | 15 % |
| Company and delivery — manufacturer vs trader, capacity, warranty and RMA, audit openness | 10 % |

| Week | Step |
| --- | --- |
| 0 | This document and Annex 1 sent to shortlisted suppliers |
| 2 | Replies received, response matrices compared |
| 3 | SDK documentation checked against our required data fields — a supplier whose SDK cannot deliver them is dropped regardless of price |
| 3–4 | Engineering samples ordered from up to three suppliers |
| 10+ | Pilot batch of 1,000 units placed with the selected supplier, scale-up to follow |
| 6–9 | Bench validation: accuracy against a reference device, battery on our duty cycle, buffer and sync, OTA from our own build, BLE traffic capture confirming no cloud connection |
| 10 | Supplier selected, factory audit, contract signed with the protocol annex |

Send to: [email] · Contact: [name, title] · Company: [company], Republic of Kazakhstan

Attached: Annex 1 — Technical Specification and Supplier Questionnaire, 98 requirements DB-nnn with priorities [M] / [P] / [O], the response matrix, the required data fields and the full regulatory list. Answer the matrix in Annex 1; use this document as the summary of what matters most.

*Republic of Kazakhstan. Confidential, issued for supplier selection only.*

# Annex 1. Screenless Health Band — Technical Specification and Supplier Questionnaire

Screen-less health monitoring band, OEM/ODM, private label

| Document | Annex 1 to the Screenless Health Band Request for Quotation dated 2 September 2026 — the complete requirement list |
| --- | --- |
| Version | 1.0 |
| Date | 2 September 2026 |
| Issued by | A healthcare IT company (Republic of Kazakhstan) |
| Product | Wrist-worn, screen-less health monitoring band |
| Pilot order | 1,000 units (tiers also quoted at 2k / 5k / 10k for the scale-up) |
| Reply deadline | 14 calendar days from receipt |

## 0. How to read and answer this document

This annex is the full requirement list. The covering RFQ document states the twelve conditions that decide the deal, the pricing table, the evaluation weights and the timeline; everything below is the detail behind them, plus the response matrix you fill in.

Every requirement carries an ID (DB-nnn) and a priority:

[M] Mandatory — a "No" here normally disqualifies the offer.

[P] Preferred — strongly wanted; a deviation must be explained and priced.

[O] Optional — quote separately, do not include in the base price.

We are not looking for a catalogue. Please answer requirement by requirement using the response matrix in Section 15 (Comply: Yes / Partial / No, plus your actual value). Suppliers who return the completed matrix, the certificate scans (Section 10) and the SDK documentation (Section 8) move directly to the sample stage. Marketing brochures without the matrix will not be evaluated.

If your closest existing model already meets most of this specification, say so and give the model number — we prefer a proven platform over a new development, provided the data-access and certification requirements are met.

## 1. Who we are and what we are building

We are a healthcare IT company from Kazakhstan. We operate a medical information system (MIS) used by state and private clinics and a patient mobile application with a nationwide user base. We are adding a hardware layer: a branded health band that supports remote monitoring of patients with chronic conditions between doctor visits — heart failure, hypertension in patients 65+, COPD/asthma and type 2 diabetes.

Key consequences for you as a manufacturer:

1. The band has no display and no stock app. All user value lives in our own iOS/Android application and in the doctor's panel inside our MIS. We will not ship your app (H Band, Da Fit, VeryFit, GloryFit, G Band or similar) to end users.

2. The data must be ours. Device → our app → our servers in Kazakhstan. No data may transit or be stored in your cloud. This is a legal requirement in our market (personal medical data localisation), not a preference.

3. Two commercial paths run in parallel and the device must fit both: - Path A — registered medical device. The band is registered in Kazakhstan as a Class 2a medical device; our company is the registration certificate holder. This requires a real quality system at the manufacturing site and a full technical file (Section 10). - Path B — wellness device, first shipment. Same hardware, sold without medical registration while Path A is in progress. We therefore need a supplier who can support the Path A documentation package without changing the hardware between paths.

## 2. Intended use and product concept

| Item | Description |
| --- | --- |
| Intended user | Adult patients, largely 50+, with a chronic condition under outpatient follow-up; also general consumers (retail channel) |
| Wear pattern | 24/7, including sleep; removed only for charging |
| Environment | Home use (IEC 60601-1-11 context), non-clinical, non-sterile |
| Primary signals | Resting and continuous heart rate, heart rate variability, SpO2, skin temperature, respiratory rate, sleep, activity, irregular-rhythm screening |
| Primary product function | Continuous background data collection + medication and measurement reminders (haptic) + one-tap confirmation by the patient |
| Explicitly out of scope | Cuffless blood pressure claims, glucose measurement claims, diagnosis on the device, any on-device machine learning |

Note on blood pressure and glucose. If your platform offers PPG-based blood pressure or glucose "trend" features, you may describe them, but they must be switchable off in firmware and must not appear on the label, packaging or in the BLE feature matrix of our production build. We will not make measurement claims we cannot register.

## 3. Scope of supply and SKUs

| SKU | Description | Status |
| --- | --- | --- |
| SKU-A (base) | Band without ECG: PPG (HR/HRV), SpO2, skin temperature, 6-axis IMU, haptics, LED, tap input | This is the SKU we intend to order. Quote as base |
| SKU-B (option) | SKU-A + single-lead ECG on demand | [O] Quote separately, incl. delta cost and certification impact |

Reason for splitting: an ECG function raises the device to a higher risk class in our registration route and lengthens approval. ECG is a second-generation SKU, but we want the price and the roadmap now, and we prefer a supplier whose platform can carry both.

Each unit shipped must include: band with fitted strap (size per order mix), charging cable or dock, retail box, printed quick-start guide in Russian and Kazakh.

## 4. Mechanical design, materials, ergonomics

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-101 | [M] | Screen-less wrist band. A status LED (1–3 colours) is acceptable and preferred; no LCD/OLED display |
| DB-102 | [M] | Total weight with strap ≤ 30 g. State exact weight |
| DB-103 | [M] | Capsule thickness ≤ 12 mm. State exact L × W × H |
| DB-104 | [M] | Quick-release, user-replaceable strap. State the mechanism and whether it is a standard lug width |
| DB-105 | [M] | Strap sizes covering wrist circumference 140–220 mm. State the size mix you offer (e.g. S/M/L) |
| DB-106 | [M] | Skin-contact materials suitable for permanent wear: medical-grade silicone or equivalent, nickel-free contact surfaces; ISO 10993 evidence required (see DB-706) |
| DB-107 | [P] | Additional strap materials (woven nylon, fabric) quoted as options |
| DB-108 | [M] | Water resistance IP68 or 5 ATM. State the test standard and depth/time actually tested |
| DB-109 | [M] | Case material and finish options (aluminium / titanium / PC+ABS). State default and premium options with price delta |
| DB-110 | [P] | Haptic (vibration) motor with programmable patterns — see DB-402. LRA preferred over ERM; state type |
| DB-111 | [M] | User input: at least a tap or capacitive touch gesture the patient can use to confirm an action (see DB-403). No mechanical button required, but state what exists |
| DB-112 | [P] | Colour options for case and strap; state minimum quantity per colour |

## 5. Sensors and measurement performance

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-201 | [M] | PPG sensor for heart rate and HRV. State: number of LEDs, wavelengths, number of photodiodes, AFE part number, sampling rate (Hz) |
| DB-202 | [M] | Continuous 24/7 HR with configurable interval (1 / 5 / 10 min), plus on-demand spot measurement triggered from our application |
| DB-203 | [M] | Beat-to-beat RR intervals available to the host application (not only an averaged BPM). State resolution (ms) |
| DB-204 | [M] | HRV metrics: at minimum RMSSD and SDNN, computed over a stated window; the underlying RR series must also be available |
| DB-205 | [M] | SpO2, continuous overnight and on demand. State method, measurement range and stated accuracy (Arms) with the validation report reference |
| DB-206 | [M] | Skin temperature sensor, resolution ≤ 0.1 °C, accuracy ≤ ±0.3 °C. State the sensor part number |
| DB-207 | [M] | 6-axis IMU (accelerometer + gyroscope). State part number and available sampling rates |
| DB-208 | [M] | Sleep detection with stages (deep / light / REM / awake) and sleep/wake timestamps |
| DB-209 | [M] | Respiratory rate during sleep |
| DB-210 | [P] | Irregular rhythm screening from PPG (episode detection, timestamped) — output must be a neutral "irregular rhythm episode" event, not a diagnosis string |
| DB-211 | [P] | Wear detection (on-wrist / off-wrist) exposed as a data field — critical for us: adherence is our main product metric |
| DB-212 | [P] | Signal-quality indicator per measurement (e.g. 0–3 or %) exposed to the host application |
| DB-213 | [O] | Single-lead ECG (SKU-B). State AFE part number, sampling rate (≥ 250 Hz required, 500 Hz preferred), electrode material, recording length, output format |
| DB-214 | [M] | Validation evidence for HR and SpO2 against a reference method, on human subjects, with number of subjects, protocol and results. State whether the report can be shared under NDA |

Accuracy targets we will verify on samples (against a certified reference device):

| Parameter | Target |
| --- | --- |
| Heart rate, resting and sleep | MAE ≤ 5 bpm or ≤ 10 %, whichever is greater |
| RR intervals, night-time | ≥ 90 % of beats detected in a clean 5-minute window |
| SpO2 | Arms ≤ 3.5 % over 70–100 % (ISO 80601-2-61 methodology) |
| Skin temperature | ± 0.3 °C vs reference |
| Sleep total time | ± 30 min vs reference on 80 % of nights |

## 6. Electronics, power, storage

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-301 | [M] | State MCU and BLE SoC exact part numbers (e.g. Nordic nRF52/nRF54, Goodix GR5515) and BLE version |
| DB-302 | [M] | Battery life ≥ 7 days with continuous 24/7 HR at 5-minute interval, overnight SpO2, sleep tracking and 3 haptic reminders per day. State the measurement conditions of your figure |
| DB-303 | [P] | Battery life ≥ 14 days under the same profile |
| DB-304 | [M] | Battery capacity (mAh), chemistry, cycle life to 80 % capacity, and the cell supplier |
| DB-305 | [M] | Magnetic charging dock or cable; full charge ≤ 2 h. Charging contacts must tolerate sweat and daily use — state the corrosion test performed |
| DB-306 | [P] | Wireless (Qi or proprietary inductive) charging as an alternative — quote separately |
| DB-307 | [M] | On-device data buffer ≥ 5 days of full-resolution data with no phone connection; ≥ 10 days preferred. State flash size and what is stored at which resolution |
| DB-308 | [M] | Data survives battery depletion and reboot until successfully synced (no silent loss). Describe the mechanism |
| DB-309 | [M] | Real-time clock with time synchronisation from the host app; all records timestamped in UTC with a stated epoch. State drift per month |
| DB-310 | [P] | Full buffered-day sync in ≤ 60 s over BLE. State the achievable throughput |
| DB-311 | [M] | Battery safety certification (IEC 62133 or equivalent) and UN 38.3 for air transport |

## 7. Firmware behaviour, reminders, OTA

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-401 | [M] | Measurement schedule (which sensor, how often, at what hours) configurable from our application at runtime, per device, without a firmware rebuild |
| DB-402 | [M] | Reminders: our app can schedule and trigger haptic + LED reminders on the device (medication, blood pressure measurement, movement). Configurable pattern, intensity and time; must fire when the phone is out of range |
| DB-403 | [M] | Patient acknowledgement: a tap/gesture on the band produces a timestamped event ("reminder acknowledged") stored in the buffer and delivered to our app. This is the core adherence mechanism of the product |
| DB-404 | [P] | Do-not-disturb window (no haptics during a configured sleep window, except critical reminders) |
| DB-405 | [M] | OTA firmware update over BLE, triggered by our application through the SDK, without your app being installed |
| DB-406 | [M] | Resumable/atomic OTA with rollback if an update fails mid-flash. Describe the recovery path |
| DB-407 | [M] | Firmware images are signed. State who holds the signing keys and whether we can hold our own keys or receive per-batch signed builds |
| DB-408 | [M] | Firmware version readable via SDK; production batches shipped with a stated, fixed firmware version. Firmware version and hardware revision must be traceable to the batch |
| DB-409 | [M] | Firmware support and security-fix policy: state the committed period in years and the response time for a critical defect |
| DB-410 | [P] | Firmware customisation available: which parameters can we change? Can we obtain source code, a build toolchain, or an escrow arrangement? |
| DB-411 | [M] | Any feature we do not license (BP, glucose, third-party cloud upload, stock-app pairing) must be absent or disabled in our production firmware |

## 8. Data access, SDK and BLE protocol — our primary selection criterion

This section decides the supplier choice. A device we cannot read fully is unusable for us at any price.

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-501 | [M] | Native BLE SDK for both iOS and Android. State minimum supported OS versions, language (Swift/Kotlin/Java), and distribution form (source, .aar/.framework, package manager) |
| DB-502 | [P] | React Native / Flutter wrapper, or a documented path to build one |
| DB-503 | [M] | SDK documentation (or a public repository link) provided before we place a sample order, under NDA if required |
| DB-504 | [M] | The SDK must expose, at minimum: per-minute or per-5-minute time series for HR, SpO2, temperature, activity; RR intervals; sleep stages; respiratory rate; battery and wear status; device events (including DB-403 acknowledgements) |
| DB-505 | [P] | Raw PPG waveform access with stated sampling rate; raw accelerometer/gyroscope access with stated rate; raw ECG waveform for SKU-B |
| DB-506 | [M] | Full offline operation: device → our app → our server, with no connection to your cloud at any point, including first-time activation and OTA. Confirm in writing |
| DB-507 | [M] | Complete BLE protocol specification (service and characteristic UUIDs, packet structures, command set, data encoding) delivered as an annex to the supply contract — our insurance against SDK discontinuation |
| DB-508 | [M] | No licence fee, per-device fee or annual fee for SDK/protocol/API access. If a fee exists, state it explicitly with the model |
| DB-509 | [M] | Protocol stability: the BLE protocol and data format remain compatible across production batches. Any breaking change requires written notice and a new SDK version supplied at no cost |
| DB-510 | [P] | A device-capability query (feature matrix readable over BLE) so our app can detect exactly what a given batch supports |
| DB-511 | [O] | If you also operate a cloud platform: REST API, webhooks/streaming, server location, and whether a deployment or data residency outside China is possible. We do not plan to use it, but we log the option |
| DB-512 | [M] | Named technical contact (engineer, not sales) for SDK integration support during our development, with a stated response SLA |

## 9. Security and privacy

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-601 | [M] | BLE pairing with bonding and link-layer encryption; the band must not deliver health data to an unbonded central |
| DB-602 | [M] | Binding of one device to one patient account, with a documented reset/unbind procedure for RMA and device reassignment |
| DB-603 | [M] | The firmware makes no network connection of its own and contains no telemetry to any third party |
| DB-604 | [P] | Resolvable private BLE address (MAC randomisation) to prevent passive tracking of the wearer |
| DB-605 | [M] | Configurable BLE advertising name and manufacturer data (we will use our own device name; it must not advertise your brand or stock app) |
| DB-606 | [M] | Buffered data on the device is erased on unbind/factory reset. Describe the procedure |
| DB-607 | [P] | Statement of known vulnerabilities and the process for handling security reports |

## 10. Regulatory, certification and technical documentation

Kazakhstan's national registration route recognises CE marking and requires evidence of a certified quality management system at the manufacturing site, plus a full technical file. Answer this section precisely — vague answers here are the most common reason we drop a supplier.

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-701 | [M] | Does the manufacturing site hold ISO 13485? Send a certificate scan showing number, notified/certification body, scope and expiry. "We can assist the customer" is not ISO 13485 |
| DB-702 | [M] | ISO 9001 certificate scan |
| DB-703 | [M] | CE marking for the proposed model: state the route — RED 2014/53/EU only, or MDR 2017/745 as a medical device — and send the Declaration of Conformity |
| DB-704 | [M] | Test reports: IEC 60601-1 (electrical safety), IEC 60601-1-2 (EMC), IEC 60601-1-11 (home healthcare environment). State which exist today, for which configuration, and which would have to be commissioned |
| DB-705 | [M] | ISO 80601-2-61 (pulse oximeter performance) report for the SpO2 function, or a clear statement that it does not exist |
| DB-706 | [M] | ISO 10993 biological evaluation for skin-contact parts — prolonged contact with intact skin: cytotoxicity, sensitisation, irritation. Send reports with the tested material identification |
| DB-707 | [M] | ISO 14971 risk management file for the product |
| DB-708 | [M] | IEC 62304 software lifecycle documentation for the device firmware, with the assigned software safety class |
| DB-709 | [M] | FCC / RoHS / REACH; battery IEC 62133 and UN 38.3 |
| DB-710 | [M] | Will you supply the full technical file for a third-country registration: design history/technical documentation, verification and validation reports, clinical evaluation or performance validation data for HR, HRV and SpO2, in English, identifying the exact tested configuration? |
| DB-711 | [M] | Regulatory role: are you willing to appear as the legal manufacturer in our dossier, or do you prefer a contract-manufacturing arrangement where we are the legal manufacturer and you are the manufacturing site? We currently plan the second option — confirm you accept it and can supply the corresponding documents |
| DB-712 | [P] | Previous experience supporting a customer's medical device registration outside China — name the market and the year (customer name may be withheld) |
| DB-713 | [M] | Are you open to an on-site or video factory audit before contract signature, and to an audit by our registration authority's assessor if required? |
| DB-714 | [P] | Traceability: UDI/serial number per unit, laser-marked or printed, readable via SDK, and a batch record retained for the warranty period |

## 11. Branding, packaging, labelling, localisation

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-801 | [M] | Full private label — our brand only. No supplier brand on device, strap, charger, packaging, manual, BLE name or firmware strings |
| DB-802 | [M] | Logo application: laser engraving on the case and/or printing on the strap. State minimum quantity and cost per method |
| DB-803 | [M] | Custom retail packaging to our artwork. Quote design/tooling and per-unit cost |
| DB-804 | [M] | Printed quick-start guide in Russian and Kazakh; we supply the translated text, you handle layout and printing. Quote per unit |
| DB-805 | [M] | Label content per our instruction, including manufacturer, model, serial/UDI, batch, date of manufacture, and — for Path A — the registration certificate holder |
| DB-806 | [M] | EAN/barcode labelling on the retail box to our numbering |
| DB-807 | [P] | Export cartons and pallet configuration stated (units per box, box dimensions, gross weight) |

## 12. Quality, reliability, acceptance

| ID | Pri | Requirement |
| --- | --- | --- |
| DB-901 | [M] | Incoming acceptance by AQL sampling: critical 0, major 0.65, minor 2.5 (ISO 2859-1). Confirm acceptance |
| DB-902 | [M] | DOA rate ≤ 0.5 % of a shipment; supplier replaces DOA units at own cost including freight |
| DB-903 | [M] | Field failure rate ≤ 3 % in the first 12 months. State your historical figure for the proposed model |
| DB-904 | [M] | Warranty: 12 months minimum, 24 months preferred, from date of sale to end user. State your terms |
| DB-905 | [M] | RMA procedure: turnaround time, who pays freight, spare-parts availability, and buffer stock policy |
| DB-906 | [M] | Reliability testing performed on the model: drop, tumble, sweat/corrosion, strap pull, charge-cycle, temperature/humidity. Send the reports or the test plan |
| DB-907 | [P] | Pre-shipment inspection by our appointed third party permitted, with production and test records made available |
| DB-908 | [M] | Firmware and hardware revision frozen for the order; any change requires our written approval and a new golden sample |

## 13. Samples, lead times, capacity

Please state, as numbers, not ranges where possible:

| ID | Question |
| --- | --- |
| DB-1001 | Cost and lead time for 3–5 engineering samples with SDK access and documentation |
| DB-1002 | Lead time for custom-branded golden samples after design freeze |
| DB-1003 | Production lead time for the pilot batch of 1,000 units, and for 10,000 units, after golden sample approval |
| DB-1004 | Monthly production capacity for this model |
| DB-1005 | Are you the manufacturer or a trading company? Factory address, year established, number of R&D engineers, number of SMT/assembly lines |
| DB-1006 | Which brands have you produced screen-less health bands for? (Names may be withheld; market and volume are sufficient) |

## 14. Commercial terms

Quote FOB (state the port) and DDP Almaty or Astana, Kazakhstan, for each tier. Prices in USD. State validity period of the quotation.

| Line item | 1,000 pcs (pilot) | 2,000 pcs | 5,000 pcs | 10,000 pcs |
| --- | --- | --- | --- | --- |
| a) Device (SKU-A), unit price |  |  |  |  |
| b) Charger / dock and cable |  |  |  |  |
| c) Custom retail packaging + printed guide (RU/KK) |  |  |  |  |
| d) One-time NRE: mould / ID change / firmware customisation |  |  |  |  |
| e) SDK, app or cloud fees, if any |  |  |  |  |
| f) Certification and test costs to be borne by us |  |  |  |  |
| g) SKU-B (ECG) delta per unit, if applicable |  |  |  |  |
| Total landed estimate, DDP Kazakhstan |  |  |  |  |

Also state:

| ID | Question |
| --- | --- |
| DB-1101 | Your standard MOQ for OEM and for ODM, separately, and whether you accept a pilot order of 1,000 units |
| DB-1102 | Payment terms (deposit %, balance, instruments accepted) |
| DB-1103 | Currency and price validity; how raw-material or exchange-rate movements are handled |
| DB-1104 | Who owns the IP of a customised design, mould and firmware developed for us? |
| DB-1105 | Are you willing to grant exclusivity for our customised model in Kazakhstan and Central Asia? On what conditions? |
| DB-1106 | Do you require a minimum annual volume to maintain the price or the exclusivity? |

Budget guidance (so neither side wastes time): our business case is built on a landed cost per unit that keeps the device affordable in a market where the device competes with consumer fitness bands. At volume, a base SKU-A quotation materially above USD 11.5 FOB at 10,000 units is unlikely to proceed, unless it is justified by certification coverage that removes cost from our registration budget. Tell us honestly where your price lands.

## 15. Supplier response matrix

Return this table completed. Add rows for every requirement you deviate from; you may return the full list, but all [M] requirements must be answered explicitly.

| Req. ID | Comply (Yes / Partial / No) | Your actual value or model detail | Comment / deviation / cost impact |
| --- | --- | --- | --- |
| DB-101 |  |  |  |
| DB-102 |  |  |  |
| … |  |  |  |

Documents to attach with your reply:

1. Completed response matrix (Section 15).

2. Completed pricing table (Section 14).

3. ISO 13485 and ISO 9001 certificate scans.

4. CE Declaration of Conformity and the test reports listed in DB-704 / DB-705.

5. ISO 10993 biocompatibility reports.

6. SDK documentation or repository link (NDA acceptable — send us your NDA and we will sign).

7. Datasheet of the proposed model, with sensor part numbers.

8. Photographs or drawings with dimensions.

Send to: [email] · Contact: [name, title] · Company: [company], Republic of Kazakhstan

We shortlist within two weeks of the reply deadline and order samples from the three strongest respondents immediately afterwards.

## Appendix A. Required data fields

The following must be retrievable through the SDK. Mark each: available now / available with firmware customisation / not available.

| # | Field | Resolution / rate expected | Status |
| --- | --- | --- | --- |
| A1 | Heart rate | 1 or 5 min continuous |  |
| A2 | Resting heart rate | daily + nightly |  |
| A3 | RR intervals | beat-to-beat, ms |  |
| A4 | HRV (RMSSD, SDNN) | nightly window |  |
| A5 | SpO2 | 5 min overnight + on demand |  |
| A6 | Respiratory rate | nightly |  |
| A7 | Skin temperature | 5–30 min, 0.1 °C |  |
| A8 | Sleep stages and sleep/wake times | per epoch |  |
| A9 | Steps, activity minutes, energy | hourly |  |
| A10 | Irregular rhythm episodes | timestamped events |  |
| A11 | Wear / off-wrist status | per interval |  |
| A12 | Signal quality indicator | per measurement |  |
| A13 | Reminder acknowledgement events (DB-403) | timestamped events |  |
| A14 | Battery level, firmware version, serial/UDI | on read |  |
| A15 | Raw PPG waveform | Hz to be stated |  |
| A16 | Raw accelerometer / gyroscope | Hz to be stated |  |
| A17 | Raw ECG (SKU-B) | ≥ 250 Hz, 500 Hz preferred |  |

## Appendix B. Evaluation criteria

Offers are scored on five weighted groups. This is how we will actually decide:

| Group | Weight | What we look at |
| --- | --- | --- |
| Data access (Section 8) | 30 % | Real SDK for both platforms, raw/RR data, offline operation, protocol annex, no fees |
| Regulatory readiness (Section 10) | 25 % | ISO 13485 at the site, existing 60601 / 80601-2-61 / 10993 reports, willingness to be the manufacturing site under our registration |
| Hardware fit (Sections 4–7) | 20 % | Battery, buffer, haptics + acknowledgement input, sensor quality, IP rating |
| Commercial (Section 14) | 15 % | Pilot price at 1k, landed cost at 10k, NRE, MOQ, payment terms, exclusivity |
| Company and delivery (Sections 12–13) | 10 % | Manufacturer vs trader, capacity, warranty and RMA, audit openness |

## Appendix C. What happens next

1. Week 0 — this document is sent to shortlisted suppliers.

2. Week 2 — replies received; response matrices compared.

3. Week 3 — SDK documentation reviewed against Appendix A; suppliers whose SDK does not deliver the required fields are dropped regardless of price.

4. Week 3–4 — engineering samples ordered from up to three suppliers.

5. Week 6–9 — bench validation on samples: accuracy against a reference device, battery under our duty cycle, buffer and sync behaviour, OTA from our own build, offline verification (BLE traffic capture confirming no cloud connection).

6. Week 10 — supplier selected, factory audit, contract with the protocol annex (DB-507).

*Republic of Kazakhstan. This document is confidential and issued for the purpose of supplier selection only.*