# Damumed Band — Request for Quotation

Screen-less health monitoring band · OEM/ODM · private label · Kazakhstan

| Document | RFQ summary — the full requirement list is Annex 1 (98 requirements, sent with this letter) |
| --- | --- |
| Version | 1.0 |
| Date | 2 September 2026 |
| Issued by | Damumed, Republic of Kazakhstan |
| Initial order | 10,000 units (quote 2k / 5k / 10k / 20k) |
| Reply deadline | 14 calendar days from receipt |

## 1. What we are buying

Damumed is a healthcare IT company from Kazakhstan. We run a medical information system used by state and private clinics and a patient mobile application with a nationwide user base. We are adding a hardware layer: a branded, screen-less wrist band for remote monitoring of patients with chronic conditions between doctor visits — heart failure, hypertension in patients 65+, COPD/asthma, type 2 diabetes — plus a retail consumer channel.

Three facts that define the whole project:

1. No display, no stock app. All value lives in our own iOS/Android application and in the doctor's panel inside our MIS. We will not ship your app (H Band, Da Fit, VeryFit, GloryFit or similar) to end users.

2. The data is ours. Device → our app → our servers in Kazakhstan. No data may transit or be stored in your cloud, at any point, including activation and firmware updates. This is a legal requirement in our market, not a preference.

3. Two paths run in parallel on the same hardware. Path A — the band is registered in Kazakhstan as a Class 2a medical device with Damumed as the certificate holder. Path B — the same device sold as a wellness product while registration is in progress. We need a supplier who can support the Path A documentation package without changing the hardware.

Two SKUs. SKU-A — no ECG: PPG (HR/HRV), SpO2, skin temperature, 6-axis IMU, haptics, LED, tap input. This is what we intend to order. SKU-B — SKU-A plus single-lead ECG: quote separately, we want the price and the roadmap now without raising the risk class today.

## 2. The twelve conditions that decide the deal

Each line references the detailed requirement in Annex 1. A "No" on any of these normally ends the evaluation, whatever the price.

| # | Condition | Annex 1 |
| --- | --- | --- |
| 1 | Native BLE SDK for both iOS and Android, documentation available before we order samples (NDA acceptable) | DB-501, DB-503 |
| 2 | Beat-to-beat RR intervals available to our app — not only an averaged BPM | DB-203 |
| 3 | SDK exposes the full data set: HR, SpO2, temperature, activity, sleep stages, respiratory rate, wear status, device events | DB-504, Annex 1 App. A |
| 4 | Fully offline operation — no connection to your cloud at any point, confirmed in writing | DB-506 |
| 5 | Complete BLE protocol specification as an annex to the supply contract — our insurance against SDK discontinuation | DB-507 |
| 6 | No licence, per-device or annual fee for SDK, protocol or API access | DB-508 |
| 7 | Haptic reminder + one-tap patient acknowledgement as a timestamped event delivered to our app — this is the core mechanism of the product, and without a display there is no other way | DB-402, DB-403 |
| 8 | Battery ≥ 7 days on our duty cycle and an on-device buffer of ≥ 5 days of full-resolution data with no phone | DB-302, DB-307 |
| 9 | ISO 13485 held by the manufacturing site, certificate scan supplied. "We can assist the customer" is not ISO 13485 | DB-701 |
| 10 | CE marking plus existing test reports: IEC 60601-1 / -1-2 / -1-11, ISO 80601-2-61 (SpO2), ISO 10993 (skin contact) — state honestly what exists and what must be commissioned | DB-703–DB-706 |
| 11 | You accept the arrangement "Damumed is the legal manufacturer, your site is the manufacturing location" and supply the corresponding technical file in English | DB-710, DB-711 |
| 12 | Full private label — Damumed only, on device, strap, charger, packaging, manual, BLE advertising name and firmware strings | DB-801 |

One more, easy to miss. If your platform offers cuffless blood pressure or glucose "trend" features, they must be switchable off in our production firmware and absent from the label, packaging and BLE feature matrix (DB-411). We will not claim what we cannot register.

Accuracy we will verify on samples against a certified reference: heart rate MAE ≤ 5 bpm (or ≤ 10 %), SpO2 Arms ≤ 3.5 % over 70–100 % per ISO 80601-2-61 methodology, skin temperature ± 0.3 °C, ≥ 90 % of beats detected in a clean 5-minute night window.

## 3. Pricing

Quote FOB (state the port) and DDP Almaty or Astana, Kazakhstan, in USD, with the validity period of the quotation.

| Line item | 2,000 pcs | 5,000 pcs | 10,000 pcs | 20,000 pcs |
| --- | --- | --- | --- | --- |
| a) Device (SKU-A), unit price |  |  |  |  |
| b) Charger / dock and cable |  |  |  |  |
| c) Custom retail packaging + printed guide (RU/KK) |  |  |  |  |
| d) One-time NRE: mould / ID change / firmware customisation |  |  |  |  |
| e) SDK, app or cloud fees, if any |  |  |  |  |
| f) Certification and test costs to be borne by us |  |  |  |  |
| g) SKU-B (ECG) delta per unit |  |  |  |  |
| Total landed estimate, DDP Kazakhstan |  |  |  |  |

Budget guidance, so neither side wastes time. Our business case is built on a landed cost that keeps a *medical* device cheaper than a consumer fitness band in our market. A base SKU-A quotation materially above USD 11.5 FOB at 10,000 units is unlikely to proceed, unless the premium is justified by certification coverage that removes cost from our registration budget. Tell us honestly where your price lands.

Also state, in one line each: your MOQ for OEM and for ODM; payment terms; IP ownership of a customised design, mould and firmware; and whether you would grant exclusivity for our customised model in Kazakhstan and Central Asia (Annex 1, DB-1101–DB-1106).

## 4. What to send us

We are not looking for a catalogue. Suppliers who return the first three items below move straight to the sample stage; brochures without the response matrix will not be evaluated.

1. Completed response matrix — Annex 1, Section 15. All [M] requirements answered explicitly.

2. Completed pricing table — Section 3 above.

3. SDK documentation or repository link (send us your NDA and we will sign it).

4. ISO 13485 and ISO 9001 certificate scans.

5. CE Declaration of Conformity and the test reports of DB-704 / DB-705.

6. ISO 10993 biocompatibility reports with the tested material identified.

7. Datasheet of the proposed model with sensor part numbers, plus photographs or drawings with dimensions.

8. Cost and lead time for 3–5 engineering samples with SDK access.

If your closest existing model already meets most of this specification, say so and give the model number. We prefer a proven platform over a new development, provided the data-access and certification requirements are met.

## 5. How we will decide, and when

| Group | Weight |
| --- | --- |
| Data access — SDK, RR/raw data, offline operation, protocol annex, no fees | 30 % |
| Regulatory readiness — ISO 13485 at the site, existing test reports, willingness to be the manufacturing site under our registration | 25 % |
| Hardware fit — battery, buffer, haptics and acknowledgement input, sensor quality, IP rating | 20 % |
| Commercial — landed cost at 10k, NRE, MOQ, payment terms, exclusivity | 15 % |
| Company and delivery — manufacturer vs trader, capacity, warranty and RMA, audit openness | 10 % |

| Week | Step |
| --- | --- |
| 0 | This document and Annex 1 sent to shortlisted suppliers |
| 2 | Replies received, response matrices compared |
| 3 | SDK documentation checked against our required data fields — a supplier whose SDK cannot deliver them is dropped regardless of price |
| 3–4 | Engineering samples ordered from up to three suppliers |
| 6–9 | Bench validation: accuracy against a reference device, battery on our duty cycle, buffer and sync, OTA from our own build, BLE traffic capture confirming no cloud connection |
| 10 | Supplier selected, factory audit, contract signed with the protocol annex |

Send to: [email] · Contact: [name, title] · Company: Damumed, Republic of Kazakhstan

Attached: Annex 1 — Technical Specification and Supplier Questionnaire, 98 requirements DB-nnn with priorities [M] / [P] / [O], the response matrix, the required data fields and the full regulatory list. Answer the matrix in Annex 1; use this document as the summary of what matters most.

*Damumed, Republic of Kazakhstan. Confidential, issued for supplier selection only.*